## Serial Numbers
LBCB BOX -Load Cell Serial Numbers
LBCB 1
141008A | X1
141029A | X2
141030A | Y1
141028A | Z1
141800A | Z2
141729A | Z3

## X1 LOAD CELL
FACTORY CALIBRATION
Load Cell Capacity | KIP | Lbs | N
| 270 | 270000 | 1201019.832
Bridge Resistance | Ω
350.5
Shunt Resistance | KΩ | Ω
| 30 | 30000
Excitation | V
10
Rated Output | Positive mV/V | Negative mV/V
| 4.1519 | -4.1445
Zero Balance | %RO
0.198
Shunt Output | Positive mV/V | Negative mV/V
| 2.9065 | -2.8988
Shunt Load Positive | KIP | Lbs | N
| 189.11 | 189110 | 841203.1868
Shunt Load Negative | KIP | Lbs | N
| -188.82 | -188820 | -839913.2025
Factory Slope | KIP/mV/V | Lbs/mV/V | N/mV/V
| 65.0885 | 65088.4721 | 289527.9475
Factory Offset | KIP | Lbs | N
| 0.5346 | 534.6 | 2378.0193
INSTALLED CALIBRATION
Shunt Resistance | KΩ | Ω
| 60 | 60000
Excitation | V
Shunt Output | Positive mV/V | Negative mV/V
| 1.4617 | -1.4578
Shunt Load Positive | KIP | Lbs | N
| 95.6738 | 95673.7752 | 425578.1532
Shunt Load Negative | KIP | Lbs | N
| -94.3525 | -94352.5292 | -419700.9583
MEASURED VOLTAGES
| V | KIP | Lbs | N
Shunt Positive | 3.0895 | 95.6738 | 95673.7752 | 425578.1532
Zero Load | 0.18767 | 0 | 0 | 0
Shunt Negative | -2.7069 | -94.3525 | -94352.5292 | -419700.9583
[5 empty] | Slope | 32.7832
[5 empty] | R2 | 0.99999
[5 empty] | Voltage @ Max Kip | 10.0661
[5 empty] | Max Kip Compression | 333.9849
[5 empty] | Max Kip Tension | -321.68
[5 empty] | Previous Slope | 32.7832

## X2 LOAD CELL
FACTORY CALIBRATION
Load Cell Capacity | KIP | Lbs | N
| 270 | 270000 | 1201019.832
Bridge Resistance | Ω
350.8
Shunt Resistance | KΩ | Ω
| 30 | 30000
Excitation | V
10
Rated Output | Positive mV/V | Negative mV/V
| 4.1245 | -4.1164
Zero Balance | %RO
0.012
Shunt Output | Positive mV/V | Negative mV/V
| 2.9028 | -2.9036
Shunt Load Positive | KIP | Lbs | N
| 190.15 | 190150 | 845829.3372
Shunt Load Negative | KIP | Lbs | N
| -190.42 | -190420 | -847030.3571
Factory Slope | KIP/mV/V | Lbs/mV/V | N/mV/V
| 65.5268 | 65526.8235 | 291477.8318
Factory Offset | KIP | Lbs | N
| 0.0324 | 32.4 | 144.1224
INSTALLED CALIBRATION
Shunt Resistance | KΩ | Ω
| 60 | 60000
Excitation | V
Shunt Output | Positive mV/V | Negative mV/V
| 1.4598 | -1.4602
Shunt Load Positive | KIP | Lbs | N
| 95.6909 | 95690.8504 | 425654.1079
Shunt Load Negative | KIP | Lbs | N
| -95.6524 | -95652.4135 | -425483.132
MEASURED VOLTAGES
| V | KIP | Lbs | N
Shunt Positive | 2.9572 | 95.6909 | 95690.8504 | 425654.1079
Zero Load | 0.082486 | 0 | 0 | 0
Shunt Negative | -2.7914 | -95.6524 | -95652.4135 | -425483.132
[5 empty] | Slope | 33.2854
[5 empty] | R2 | 1
[5 empty] | Voltage @ Max Kip | 9.9142
[5 empty] | Max Kip Compression | 335.5998
[5 empty] | Max Kip Tension | -330.1087
[5 empty] | Previous Slope | 33.2854

## Y1 LOAD CELL
FACTORY CALIBRATION
Load Cell Capacity | KIP | Lbs | N
| 270 | 270000 | 1201019.832
Bridge Resistance | Ω
352.5
Shunt Resistance | KΩ | Ω
| 30 | 30000
Excitation | V
10
Rated Output | Positive mV/V | Negative mV/V
| 4.1201 | -4.1105
Zero Balance | %RO
0.07
Shunt Output | Positive mV/V | Negative mV/V
| 2.9152 | -2.8994
Shunt Load Positive | KIP | Lbs | N
| 191.19 | 191190 | 850455.4877
Shunt Load Negative | KIP | Lbs | N
| -190.39 | -190390 | -846896.9104
Factory Slope | KIP/mV/V | Lbs/mV/V | N/mV/V
| 65.6088 | 65608.8256 | 291842.5952
Factory Offset | KIP | Lbs | N
| 0.189 | 189 | 840.7139
INSTALLED CALIBRATION
Shunt Resistance | KΩ | Ω
| 60 | 60000
Excitation | V
Shunt Output | Positive mV/V | Negative mV/V
| 1.4661 | -1.4582
Shunt Load Positive | KIP | Lbs | N
| 96.379 | 96378.9773 | 428715.0487
Shunt Load Negative | KIP | Lbs | N
| -95.4796 | -95479.6403 | -424714.5984
MEASURED VOLTAGES
| V | KIP | Lbs | N
Shunt Positive | 3.0139 | 96.379 | 96378.9773 | 428715.0487
Zero Load | 0.12667 | 0 | 0 | 0
Shunt Negative | -2.745 | -95.4796 | -95479.6403 | -424714.5984
[5 empty] | Slope | 33.3153
[5 empty] | R2 | 1
[5 empty] | Voltage @ Max Kip | 9.9054
[5 empty] | Max Kip Compression | 337.3729
[5 empty] | Max Kip Tension | -328.9325
[5 empty] | Previous Slope | 33.3153

## Z1 LOAD CELL
FACTORY CALIBRATION
Load Cell Capacity | KIP | Lbs | N
| 270 | 270000 | 1201019.832
Bridge Resistance | Ω
351.1
Shunt Resistance | KΩ | Ω
| 30 | 30000
Excitation | V
10
Rated Output | Positive mV/V | Negative mV/V
| 4.1167 | -4.1089
Zero Balance | %RO
-0.041
Shunt Output | Positive mV/V | Negative mV/V
| 2.903 | -2.9024
Shunt Load Positive | KIP | Lbs | N
| 190.53 | 190530 | 847519.6614
Shunt Load Negative | KIP | Lbs | N
| -190.69 | -190690 | -848231.3769
Factory Slope | KIP/mV/V | Lbs/mV/V | N/mV/V
| 65.6487 | 65648.7065 | 292019.9942
Factory Offset | KIP | Lbs | N
| -0.1107 | -110.7 | -492.4181
INSTALLED CALIBRATION
Shunt Resistance | KΩ | Ω
| 60 | 60000
Excitation | V
Shunt Output | Positive mV/V | Negative mV/V
| 1.4599 | -1.4596
Shunt Load Positive | KIP | Lbs | N
| 95.7328 | 95732.7536 | 425840.5023
Shunt Load Negative | KIP | Lbs | N
| -95.9343 | -95934.3444 | -426737.2229
MEASURED VOLTAGES
| V | KIP | Lbs | N
Shunt Positive | 3.603 | 95.7328 | 95732.7536 | 425840.5023
Zero Load | 0.66711 | 0 | 0 | 0
Shunt Negative | -2.2703 | -95.9343 | -95934.3444 | -426737.2229
[5 empty] | Slope | 32.6334
[5 empty] | R2 | 1
[5 empty] | Voltage @ Max Kip | 10.1123
[5 empty] | Max Kip Compression | 348.1043
[5 empty] | Max Kip Tension | -304.5639
[5 empty] | Previous Slope | 32.6334

## Z2 LOAD CELL
FACTORY CALIBRATION
Load Cell Capacity | KIP | Lbs | N
| 270 | 270000 | 1201019.832
Bridge Resistance | Ω
350.8
Shunt Resistance | KΩ | Ω
| 30 | 30000
Excitation | V
10
Rated Output | Positive mV/V | Negative mV/V
| 4.1248 | -4.108
Zero Balance | %RO
0.002
Shunt Output | Positive mV/V | Negative mV/V
| 2.9061 | -2.9034
Shunt Load Positive | KIP | Lbs | N
| 190.36 | 190360 | 846763.4638
Shunt Load Negative | KIP | Lbs | N
| -190.75 | -190750 | -848498.2702
Factory Slope | KIP/mV/V | Lbs/mV/V | N/mV/V
| 65.5913 | 65591.2934 | 291764.6079
Factory Offset | KIP | Lbs | N
| 0.0054 | 5.4 | 24.0204
INSTALLED CALIBRATION
Shunt Resistance | KΩ | Ω
| 60 | 60000
Excitation | V
Shunt Output | Positive mV/V | Negative mV/V
| 1.4615 | -1.4601
Shunt Load Positive | KIP | Lbs | N
| 95.8668 | 95866.8206 | 426436.862
Shunt Load Negative | KIP | Lbs | N
| -95.767 | -95766.9576 | -425992.6495
MEASURED VOLTAGES
| V | KIP | Lbs | N
Shunt Positive | 2.8842 | 95.8668 | 95866.8206 | 426436.862
Zero Load | -0.023525 | 0 | 0 | 0
Shunt Negative | -2.9274 | -95.767 | -95766.9576 | -425992.6495
[5 empty] | Slope | 32.9745
[5 empty] | R2 | 1
[5 empty] | Voltage @ Max Kip | 10.0077
[5 empty] | Max Kip Compression | 328.969
[5 empty] | Max Kip Tension | -330.5204
[5 empty] | Previous Slope | 32.9745
5.7807
-0.023525
-5.8607

## Z3 LOAD CELL
FACTORY CALIBRATION
Load Cell Capacity | KIP | Lbs | N
| 270 | 270000 | 1201019.832
Bridge Resistance | Ω
350.8
Shunt Resistance | KΩ | Ω
| 30 | 30000
Excitation | V
10
Rated Output | Positive mV/V | Negative mV/V
| 4.1295 | -4.1234
Zero Balance | %RO
-0.172
Shunt Output | Positive mV/V | Negative mV/V
| 2.9033 | -2.9026
Shunt Load Positive | KIP | Lbs | N
| 189.94 | 189940 | 844895.2107
Shunt Load Negative | KIP | Lbs | N
| -190.03 | -190030 | -845295.5506
Factory Slope | KIP/mV/V | Lbs/mV/V | N/mV/V
| 65.4315 | 65431.5453 | 291054.013
Factory Offset | KIP | Lbs | N
| -0.4644 | -464.4 | -2065.7541
INSTALLED CALIBRATION
Shunt Resistance | KΩ | Ω
| 60 | 60000
Excitation | V
Shunt Output | Positive mV/V | Negative mV/V
| 1.4601 | -1.4597
Shunt Load Positive | KIP | Lbs | N
| 95.0714 | 95071.4127 | 422898.7117
Shunt Load Negative | KIP | Lbs | N
| -95.9772 | -95977.1786 | -426927.7589
MEASURED VOLTAGES
| V | KIP | Lbs | N
Shunt Positive | 2.9837 | 95.0714 | 95071.4127 | 422898.7117
Zero Load | 0.069251 | 0 | 0 | 0
Shunt Negative | -2.8437 | -95.9772 | -95977.1786 | -426927.7589
[5 empty] | Slope | 32.7849
[5 empty] | R2 | 0.99999
[5 empty] | Voltage @ Max Kip | 10.0656
[5 empty] | Max Kip Compression | 330.1195
[5 empty] | Max Kip Tension | -325.5788
5.9284
0.069251 |  |  |  |  | Previous Slope | 32.7849
-5.6589