## Serial Numbers
LBCB BOX -Load Cell Serial Numbers
LBCB 2 (I Think)
141727A | X1 | 33.2718
141010A | X2 | 32.6183
141799A | Y1 | 33.1285
141805A | Z1 | 32.9657
141804A | Z2 | 33.2995
141009A | Z3 | 32.948

## X1 LOAD CELL
FACTORY CALIBRATION
Load Cell Capacity | KIP | Lbs | N
| 270 | 270000 | 1201019.832
Bridge Resistance | Ω
350.4
Shunt Resistance | KΩ | Ω
| 30 | 30000
Excitation | V
10
Rated Output | Positive mV/V | Negative mV/V
| 4.1387 | -4.1226
Zero Balance | %RO
0.131
Shunt Output | Positive mV/V | Negative mV/V
| 2.8991 | -2.8988
Shunt Load Positive | KIP | Lbs | N
| 189.26 | 189260 | 841870.42
Shunt Load Negative | KIP | Lbs | N
| -189.78 | -189780 | -844183.4952
Factory Slope | KIP/mV/V | Lbs/mV/V | N/mV/V
| 65.365 | 65365.0152 | 290758.0725
Factory Offset | KIP | Lbs | N
| 0.3537 | 353.7 | 1573.336
INSTALLED CALIBRATION
Shunt Resistance | KΩ | Ω
| 60 | 60000
Excitation | V
Shunt Output | Positive mV/V | Negative mV/V
| 1.458 | -1.4578
Shunt Load Positive | KIP | Lbs | N
| 95.6537 | 95653.6842 | 425488.7842
Shunt Load Negative | KIP | Lbs | N
| -94.9364 | -94936.4225 | -422298.2453
MEASURED VOLTAGES
| V | KIP | Lbs | N
Shunt Positive | 3.0278 | 95.6537 | 95653.6842 | 425488.7842
Zero Load | 0.16358 | 0 | 0 | 0
Shunt Negative | -2.7005 | -94.9364 | -94936.4225 | -422298.2453
[5 empty] | Slope | 33.2718
[5 empty] | R2 | 1
[5 empty] | Voltage @ Max Kip | 9.9183
[5 empty] | Max Kip Compression | 338.1612
[5 empty] | Max Kip Tension | -327.2757
[5 empty] | Previous Slope | 33.2544

## X2 LOAD CELL
FACTORY CALIBRATION
Load Cell Capacity | KIP | Lbs | N
| 270 | 270000 | 1201019.832
Bridge Resistance | Ω
350.6
Shunt Resistance | KΩ | Ω
| 30 | 30000
Excitation | V
10
Rated Output | Positive mV/V | Negative mV/V
| 4.1667 | -4.1502
Zero Balance | %RO
-0.098
Shunt Output | Positive mV/V | Negative mV/V
| 2.9036 | -2.8873
Shunt Load Positive | KIP | Lbs | N
| 188.32 | 188320 | 837689.0917
Shunt Load Negative | KIP | Lbs | N
| -187.8 | -187800 | -835376.0165
Factory Slope | KIP/mV/V | Lbs/mV/V | N/mV/V
| 64.928 | 64928.0381 | 288814.3015
Factory Offset | KIP | Lbs | N
| -0.2646 | -264.6 | -1176.9994
INSTALLED CALIBRATION
Shunt Resistance | KΩ | Ω
| 60 | 60000
Excitation | V
Shunt Output | Positive mV/V | Negative mV/V
| 1.4602 | -1.452
Shunt Load Positive | KIP | Lbs | N
| 94.5455 | 94545.5332 | 420559.483
Shunt Load Negative | KIP | Lbs | N
| -94.5425 | -94542.4956 | -420545.971
MEASURED VOLTAGES
| V | KIP | Lbs | N
Shunt Positive | 3.0483 | 94.5455 | 94545.5332 | 420559.483
Zero Load | 0.14117 | 0 | 0 | 0
Shunt Negative | -2.7486 | -94.5425 | -94542.4956 | -420545.971
[5 empty] | Slope | 32.6183
[5 empty] | R2 | 1
[5 empty] | Voltage @ Max Kip | 10.117
[5 empty] | Max Kip Compression | 330.7879
[5 empty] | Max Kip Tension | -321.5786
[5 empty] | Previous Slope | 32.6078

## Y1 LOAD CELL
FACTORY CALIBRATION
Load Cell Capacity | KIP | Lbs | N
| 270 | 270000 | 1201019.832
Bridge Resistance | Ω
350.7
Shunt Resistance | KΩ | Ω
| 30 | 30000
Excitation | V
10
Rated Output | Positive mV/V | Negative mV/V
| 4.1011 | -4.0886
Zero Balance | %RO
0.112
Shunt Output | Positive mV/V | Negative mV/V
| 2.9046 | -2.9065
Shunt Load Positive | KIP | Lbs | N
| 191.36 | 191360 | 851211.6854
Shunt Load Negative | KIP | Lbs | N
| -191.9 | -191900 | -853613.725
Factory Slope | KIP/mV/V | Lbs/mV/V | N/mV/V
| 65.9365 | 65936.4812 | 293300.0799
Factory Offset | KIP | Lbs | N
| 0.3024 | 302.4 | 1345.1422
INSTALLED CALIBRATION
Shunt Resistance | KΩ | Ω
| 60 | 60000
Excitation | V
Shunt Output | Positive mV/V | Negative mV/V
| 1.4607 | -1.4617
Shunt Load Positive | KIP | Lbs | N
| 96.6184 | 96618.4137 | 429780.1147
Shunt Load Negative | KIP | Lbs | N
| -96.0766 | -96076.6173 | -427370.0845
MEASURED VOLTAGES
| V | KIP | Lbs | N
Shunt Positive | 3.0439 | 96.6184 | 96618.4137 | 429780.1147
Zero Load | 0.1365 | 0 | 0 | 0
Shunt Negative | -2.7727 | -96.0766 | -96076.6173 | -427370.0845
[5 empty] | Slope | 33.1285
[5 empty] | R2 | 1
[5 empty] | Voltage @ Max Kip | 9.9612
[5 empty] | Max Kip Compression | 335.8066
[5 empty] | Max Kip Tension | -326.7625
[5 empty] | Previous Slope | 33.114

## Z1 LOAD CELL
FACTORY CALIBRATION
Load Cell Capacity | KIP | Lbs | N
| 270 | 270000 | 1201019.832
Bridge Resistance | Ω
350.5
Shunt Resistance | KΩ | Ω
| 30 | 30000
Excitation | V
10
Rated Output | Positive mV/V | Negative mV/V
| 4.1441 | -4.1338
Zero Balance | %RO
0.273
Shunt Output | Positive mV/V | Negative mV/V
| 2.8953 | -2.9044
Shunt Load Positive | KIP | Lbs | N
| 188.77 | 188770 | 839690.7914
Shunt Load Negative | KIP | Lbs | N
| -189.67 | -189670 | -843694.1909
Factory Slope | KIP/mV/V | Lbs/mV/V | N/mV/V
| 65.2339 | 65233.9361 | 290175.0038
Factory Offset | KIP | Lbs | N
| 0.7371 | 737.1 | 3278.7841
INSTALLED CALIBRATION
Shunt Resistance | KΩ | Ω
| 60 | 60000
Excitation | V
Shunt Output | Positive mV/V | Negative mV/V
| 1.4561 | -1.4606
Shunt Load Positive | KIP | Lbs | N
| 95.7215 | 95721.4668 | 425790.2964
Shunt Load Negative | KIP | Lbs | N
| -94.5458 | -94545.8051 | -420560.6923
MEASURED VOLTAGES
| V | KIP | Lbs | N
Shunt Positive | 3.3293 | 95.7215 | 95721.4668 | 425790.2964
Zero Load | 0.44909 | 0 | 0 | 0
Shunt Negative | -2.4423 | -94.5458 | -94545.8051 | -420560.6923
[5 empty] | Slope | 32.9657
[5 empty] | R2 | 0.99998
[5 empty] | Voltage @ Max Kip | 10.0104
[5 empty] | Max Kip Compression | 344.4613
[5 empty] | Max Kip Tension | -314.852
[5 empty] | Previous Slope | 32.9653

## Z2 LOAD CELL
FACTORY CALIBRATION
Load Cell Capacity | KIP | Lbs | N
| 270 | 270000 | 1201019.832
Bridge Resistance | Ω
350.7
Shunt Resistance | KΩ | Ω
| 30 | 30000
Excitation | V
10
Rated Output | Positive mV/V | Negative mV/V
| 4.1199 | -4.112
Zero Balance | %RO
-0.177
Shunt Output | Positive mV/V | Negative mV/V
| 2.9026 | -2.9064
Shunt Load Positive | KIP | Lbs | N
| 190.3 | 190300 | 846496.5705
Shunt Load Negative | KIP | Lbs | N
| -190.8 | -190800 | -848720.6813
Factory Slope | KIP/mV/V | Lbs/mV/V | N/mV/V
| 65.5985 | 65598.4645 | 291796.5068
Factory Offset | KIP | Lbs | N
| -0.4779 | -477.9 | -2125.8051
INSTALLED CALIBRATION
Shunt Resistance | KΩ | Ω
| 60 | 60000
Excitation | V
Shunt Output | Positive mV/V | Negative mV/V
| 1.4597 | -1.4616
Shunt Load Positive | KIP | Lbs | N
| 95.2784 | 95278.3798 | 423819.3469
Shunt Load Negative | KIP | Lbs | N
| -96.3595 | -96359.5411 | -428628.5921
MEASURED VOLTAGES
| V | KIP | Lbs | N
Shunt Positive | 3.2494 | 95.2784 | 95278.3798 | 423819.3469
Zero Load | 0.37279 | 0 | 0 | 0
Shunt Negative | -2.5056 | -96.3595 | -96359.5411 | -428628.5921
[5 empty] | Slope | 33.2995
[5 empty] | R2 | 0.99999
[5 empty] | Voltage @ Max Kip | 9.9101
[5 empty] | Max Kip Compression | 345.4089
[5 empty] | Max Kip Tension | -320.5813
[5 empty] | Previous Slope | 33.2916

## Z3 LOAD CELL
FACTORY CALIBRATION
Load Cell Capacity | KIP | Lbs | N
| 270 | 270000 | 1201019.832
Bridge Resistance | Ω
350.8
Shunt Resistance | KΩ | Ω
| 30 | 30000
Excitation | V
10
Rated Output | Positive mV/V | Negative mV/V
| 4.1395 | -4.1292
Zero Balance | %RO
-0.172
Shunt Output | Positive mV/V | Negative mV/V
| 2.9005 | -2.9019
Shunt Load Positive | KIP | Lbs | N
| 189.33 | 189330 | 842181.7955
Shunt Load Negative | KIP | Lbs | N
| -189.69 | -189690 | -843783.1553
Factory Slope | KIP/mV/V | Lbs/mV/V | N/mV/V
| 65.3065 | 65306.5173 | 290497.8611
Factory Offset | KIP | Lbs | N
| -0.4644 | -464.4 | -2065.7541
INSTALLED CALIBRATION
Shunt Resistance | KΩ | Ω
| 60 | 60000
Excitation | V
Shunt Output | Positive mV/V | Negative mV/V
| 1.4587 | -1.4594
Shunt Load Positive | KIP | Lbs | N
| 94.7969 | 94796.9004 | 421677.62
Shunt Load Negative | KIP | Lbs | N
| -95.7717 | -95771.6807 | -426013.6587
MEASURED VOLTAGES
| V | KIP | Lbs | N
Shunt Positive | 3.19 | 94.7969 | 94796.9004 | 421677.62
Zero Load | 0.29833 | 0 | 0 | 0
Shunt Negative | -2.594 | -95.7717 | -95771.6807 | -426013.6587
[5 empty] | Slope | 32.948
[5 empty] | R2 | 0.99999
[5 empty] | Voltage @ Max Kip | 10.0158
[5 empty] | Max Kip Compression | 339.3095
[5 empty] | Max Kip Tension | -319.6505
[5 empty] | Previous Slope | 32.9414