## Startup
Test | Procedure | Result | Passed | Failed | Notes | 2012-03-21
Real Startup | Open up the OM Latest Build folder and double click on the "LBCB Operation Manager" executable | Two widgets should be running. For the Mixed Mode widget, target and input mode window should be grayed out.
| Press Start Autobalancing on the LBCB Operation Manager | The auto balance initialization and Command Displacement Difference window should appear
| Select autobalance on all auto balance windows that appear. | A warning popup should show up. After ok'ing the popup, the system should start autobalancing. After it is balanced there will most likely be numerous limit faults because the fake initial position is deliberately skewed. The limits will need to be adjusted to correct these faults before any further testing is done.
| Press control on one or both LBCBs | The Mixed Mode Widget should now allow inputs
Turn On Oil | Check the difference between command and displacement for all of the actuators and DOFs. | The differences should be insignificant. If they are not the LBCB will jump to its initial position once oil pressure is applied
| For the small LBCBs turn low pressure on. | Verify that the forces have not changed significantly
| For the large LBCBs turn low pressure on and unlock the actuators one at a time starting with the Z actuators. | Verify that the forces have not changed significantly
| Switch to high pressure | The forces will jump somewhat but then settle to an equilibrium. This equilibrium is zero unless the LBCB is connected to a specimen. If a specimen is connected then the equilbrium will be close to the last position the LBCB was moved to before it was shut down.

## OM Setup Test
Test | Procedure | Result | Passed | Failed | Notes | 2012-03-21
OM Setup Test | In the OM Latest Build folder, make a copy of the Config folder so that the configuration can be restored at the end of the test
| Open up the OM Latest Build folder and double click on the "OM Setup" executable | OM Setup should display with the Load button blinking
| Select 1/5 scale two box and press Load | OM Setup will show the saved values for the Actuator Limits.
| Change all of the limit values to unique values | The Save button should be blinking
Switch to LBCB2 and enter unique values for each limit.
| Press Save and then OK | OM Setup should exit
| Restart the OM Setup program and load the 1/5 scale 2 box configuration | Verify that all of the limit values for both LBCBs are the same as what was entered.
| Repeat the value entry tests for Cartesian Limits | Verify that all of the limit values for both LBCBs are the same as what was entered.
| Repeat the value entry tests for Autobalance | Verify that all of the values for both LBCBs are the same as what was entered.
| Repeat the value entry tests for Timing and Logging | Verify that all of the values are the same as what was entered. |  |  | The values are near what was entered, but not exact (ex. 38 to 41.6667). Cannot switch between LBCB 1 and LBCB 2. Looking at all OM Locations, only "Crane Bay One LBCB" and "Portable LBCB" switch between LBCB 1 and LBCB 2.
| Repeat the value entry tests for LBCB Position | Verify that all of the values for both LBCBs are the same as what was entered.
| Repeat the value entry tests for Mixed Mode | Verify that all of the values for both LBCBs are the same as what was entered.
| Repeat the value entry tests for Calibration | Verify that all of the values for both LBCBs are the same as what was entered.
| Repeat the value entry tests for Channel Names | Verify that all of the values for both LBCBs are the same as what was entered.
| Repeat the value entry tests for External Sensors | Verify that all of the values for both LBCBs are the same as what was entered. |  |  | Do LBCB 1 and LBCB 2 have the same values? i.e. changing one changes the other
| Set Rigid Connection to "Yes". Save the change. | Verify that LBCB2 cannot be selected for the Mixed Mode, and Cartesian Limits pages. |  |  | LBCB 2 cannot be selected for Timing and Logging
| Restart the OM Setup program and load the 1/5 scale 2 box configuration | Verify that the Rigid Connection is still "Yes".
Operation Manager Configuration Test | Open up the OM Latest Build folder and double click on the "LBCB Operation Manager" executable | Two widgets should be running. For the Mixed Mode widget, target and input mode window should be grayed out. |  |  | POP-UP MESSAGE: DAQ Receive Channels not configured correctly: ,,,,,,,,,,,LPot3,LPot4,LPot1,LPot2 (swith "Real" to "FAKE!!!")
| Select the Edit->Position menu item | The LBCB Position widget should appear
Change all of the position values to unique values and press OK. Do this for both LBCBs
| Reselect Edit->Position | Verify that all of the values are still changed
Press OK
| Restart the LBCB Operation Manager program. Select Edit->Position again | Verify that all of the values have returned to their original values.
Change all of the position values to unique values and press OK. Do this for both LBCBs
Select File->Save and restart the program
| Select Edit->Position again | Verify that all of the values are still changed
| Repeat this test with Edit->Limits. Note There are 16 different displays when selecting a limit set and an LBCB | Verify that the values change when they are supposed to. |  |  | Cannot switch from LBCB 1 to LBCB 2 on Limits tab. And LBCB on Limits tab does not control LBCB on main OM
| Repeat the test for Edit->Middle Loop | Verify that the values change when they are supposed to.
| Repeat the test for Edit->Mixed Mode Loop | Verify that the values change when they are supposed to. |  |  | POP-UP MESSAGE: Selection "Mixed-mode Loop" not recognized!
| Repeat the test for Edit->Timing & Logging | Verify that the values change when they are supposed to.
| Repeat the test for the "Auto Balance Control" and "Auto Balance Error Window" in the Autobalance tab | Verify that the values change when they are supposed to.
| Repeat the test for the stiffness matrix, Jacobian Gain, Max Increment, Force Group Priorities, Force Error Range, Stiffness Limits, and Force Control Parameters in the Mixed Mode tab | Verify that the values change when they are supposed to.

## No Oil Test
Test | Procedure | Result | Passed | Failed | Notes | 2012-03-21
Start | Real Startup procedure needs to done up to the point where the control for each LBCB is turned on. Oil pressure should not be used for this test. | The Mixed Mode widget should be ready to accept commands.
Target Patterns | Set Target Pattern to Clear Targets and press Apply | No Targets should be displayed
| Set Target Pattern to Reset to Zero and press Apply | All displacements and rotations for all LBCBs should be displayed in the Targets window with values of zero
| Set Target Pattern to Displacements & Rotations and press Apply | All displacements and rotations for all LBCBs should be displayed in the Targets window with their current value
| Set Target Pattern to Displacements Only and press Apply | All displacements for all LBCBs should be displayed in the Targets window with their current value
| Set Target Pattern to Rotations Only and press Apply | All rotations for all LBCBs should be displayed in the Targets window with their current value
| Set Target Pattern to Forces & Moments and press Apply | All forces and moments for all LBCBs should be displayed in the Targets window with their current value
| Set Target Pattern to Forces Only and press Apply | All forces for all LBCBs should be displayed in the Targets window with their current value
| Set Target Pattern to Moments Only and press Apply | All moments for all LBCBs should be displayed in the Targets window with their current value
Manual input complete step | Input a target of displacement Dx within the displacement limits and press "Set Target" | Start button should blink
Input a ramp time of 10 seconds
| Press start. | Mixed Mode State display should iterate through 'Ramp' and then 'Stab' and 'Hold' for Displacement and for Force. The command at the end of the ramp should be the same as the target. The MM State should be 'Idle' at the end of the ramp.
| Select 'Reset Targets To Zero' target pattern and click on 'Apply' | Target should consist of all 6 displacement DOFs set to zero. 'Start' button should blink.
| Press start. | Mixed Mode State display should iterate through 'Ramp', 'Stab', and 'Hold'. Command at the end of the ramp should be all zeros. The MM State should be 'Idle' at the end of the ramp.
Repeat test for Dy
Repeat test for Dz
Repeat test for Rx
Repeat test for Ry
Repeat test for Rz
Repeat test by inputing a combintion of DOFs. For example use Dx and Ry.
Repeat test with LBCB 2 if testing with the two box setup
| Repeat test with the rigid connection button set to 'YES' if testing with the two box setup. This needs to be done in OM Setup. After imposing "rigid connection" to "Yes", for all target patterns when LBCB2 is displayed, hitting "Set Target" grays out the LBCB2 values and gives the error message "Target Type Detected". "Set Target" must be hit again to execute the targets. Subsequent tests do not produce the error, only initially. | In this case LBCB 2 becomes a slave to LBCB 1 and only LBCB 1 targets are possible
Repeat test for the target patterns "Displacements Only", "Rotations Only", and "Displacements & Rotations"
Manual input with 'Stop' button | Input a target of displacement Dx within the displacement limits and press "Set Target" | Start button should blink
Input a ramp time of 10 seconds
| Press start. | Mixed Mode State display should indicate the start of a ramp.
| Press stop while the MM State is at 'Ramping' | MM State should be 'Idle'. The current position should not be the target
| Press Set Target. Do not change the value of the target | Start button should blink
| Press Start | Ramp should resume from the current position to the original target. Mixed Mode State display should iterate through 'Ramp', 'Stab', and 'Hold'. The MM State should be 'Idle' at the end of the ramp.
| Select 'Reset Targets To Zero' target pattern and click on 'Apply' | Target should consist of all 6 displacement DOFs set to zero.
| Press 'Set Target' | 'Start' button should blink.
| Press start. | Mixed Mode State display should iterate through 'Ramp'.
| Press stop during the hold state | The Mixed Mode State display should be 'Idle'
Manual Input with a limit fault | Input a target displacement for Dx beyond the Dx limit and press "Set Target". | Start button should blink
| Press start | The ramp should run until the DOF reaches its limit. At this point the MM state display should be 'Idle', The limit fault light should be red, and a message should be displayed that indicates which DOF has faulted.
| Open the limits display and extend the limit so that the fault is cleared. | The fault light should turn green. The fault message should clear after about 5 seconds. The command for the faulted DOF should not change significantly.
Repeat test for Dy
Repeat test for Dz
Repeat test for Rx
Repeat test for Ry
Repeat test for Rz
Repeat test using the actuator displacement limits for all six actuators
Repeat test with LBCB 2 if testing with the two box setup
| Repeat test with the rigid connection button set to 'YES' if testing with the two box setup | In this case LBCB 2 limits are no longer available and the LBCB 2 movement is defined by the LBCB 1 limits
Input File Run/Pause/Stop/Reset/Fault | Select Input File Mode
| Press Load and select an input file for each LBCB being tested that is within the limits of the DOFs being controlled for the first few steps. Later on in the test there needs to be a target that exceeds a limit. For Example, if testing on the large scale system, select the file "Large Scale 6 DOF Large Steps" for each LBCB. For this file limits set to be exceeded during the test can be Dz for LBCB1 and Rz for LBCB2 | Pause/Running should be paused.
| Press Run | The accept button should blink. Pause/Running should be running.
| Press Accept | The start button should blink. The accept button stops blinking.
| Press start. | The start button stops blinking. The ramp for the first target should complete. The accept button should blink and the next target should be displayed when the ramp is complete.
| Press Decline | Pause/Running should be paused. The accept button stops blinking.
| Press Run | The accept button should blink. Pause/Running should be running.
| Press Accept | The start button should blink.
| Press Stop. | Pause/Running should be paused. The start button should stop blinking
| Press Run | Pause/Running should be running. The accept button should start blinking.
| Switch Propose:Manual to Auto and Execute:Manual to Auto when the Accept and Start buttons are blinking. | Multiple targets are executed until a target exceeds a limit and a fault is generated. The fault light should become red and a message indicating the DOF that has faulted should be displayed. Pause/Running should be paused.
| Open the Limits display and check the Local limit for a DOF. Record the global limit also. Close the limits display and set a target that is just beyond the largest target in the input file. | The fault light should turn green. The fault message should clear. The command for the faulted DOF should not change significantly.
| Press Run. Switch Propose:Manual to Auto and Execute:Manual to Auto | The rest of the targets should be executed. After the input file is done, paused/running should be at paused. Make sure that all of the targets were executed.
| Press Reset and then press Run | The input file should start from the beginning.
| Press Stop | Pause/Running should be paused.
| Press Run | Test execution should continue from the current target.
NTCP | Move the LBCBs to a position close to the nominal origin but not to exactly zero
| Set the offsets and switch from Global to Local coordinates | All of the displacments should read zero
| Switch to NTCP mode and press "Connect NTCP" | The Socket widget should appear and be running
Open up a command window. Change to the "OM Latest Build" folder located at "C:\Documents and Settings\All Users\Desktop\OM Latest Build"
| Type "testNtcpLink.pl --lbcb=?" in the command window where ? Is the number of LBCBs you are controlling | The script should start running. The Mixed Mode window will have a blinking Accept button. The Targets window will have a target of 1 for the DX of each LBCB. The test stats window should display step "1 0 4"
| Press the Accept button | The Start button should blink
| Press the start button | The step should run. After it is completed the next step is sent from the script. The (target) Dx should be at -1. The Accept button should be blinking.
| Press the Decline button | The command window should display "ABORTING PROPOSE COMMAND". The script will immediately send another step with Dx = 1. The Accept button should be blinking. During an actual test the LbcbPlugin will pause running when it receives a "decline" response.
| Press the Accept button and then the Stop button | The command window should display "ABORTING EXECUTE COMMAND". The script will immediately send another step with Dx = -1. The Accept button should be blinking. During an actual test the LbcbPlugin will pause running when it receives a "stop" response.
| Set the limit for Dx to something more than -1. Press Accept and then Start | The ramp should run until the limit is reached. The command window should display "ABORTING EXECUTE COMMAND". The script should send another step with Dx = 1. The Accept button should be blinking.
| Press "Disconnect NTCP" and type <CTRL> C in the command window | The Socket widget should disappear. The Script should stop running.
| Press "Connect NTCP" to restart the widget for each of the following tests | The Socket widget should appear and be running
| Type "testNtcpLink.pl --lbcb=? --option=1" in the command window where ? Is the number of LBCBs you are controlling | The script should start running. The Mixed Mode window will have a blinking Accept button. The Targets window will have a target of 1 for the DX of each LBCB. The test stats window should display step "1 0 4"
| Press Accept and then Start. While the ramp is running switch both Execute and Propose buttons to Auto. | The script will execute three steps. At the fourth step the script will send a close-session command to the OM after the propose message. At this point NTCP should disconnect and the Socket widget will disappear
| Type "testNtcpLink.pl --lbcb=? --option=2" in the command window where ? Is the number of LBCBs you are controlling | The script should start running. The Mixed Mode window will have a blinking Accept button. The Targets window will have a target of 1 for the DX of each LBCB. The test stats window should display step "1 0 4"
| Press Accept and then Start. While the ramp is running switch both Execute and Propose buttons to Auto. | The script will execute three steps. At the fourth step the script will send a close-session command to the OM after the OM accepts the propose message. At this point NTCP should disconnect and the Socket widget will disappear
| Type "testNtcpLink.pl --lbcb=? --option=3" in the command window where ? Is the number of LBCBs you are controlling | The script should start running. The Mixed Mode window will have a blinking Accept button. The Targets window will have a target of 1 for the DX of each LBCB. The test stats window should display step "1 0 4"
| Press Accept and then Start. While the ramp is running switch both Execute and Propose buttons to Auto. | The script will execute three steps. At the fourth step the script will send a close-session command to the OM after the execute message. At this point NTCP should disconnect and the Socket widget will disappear
| Type "testNtcpLink.pl --lbcb=? --option=4" in the command window where ? Is the number of LBCBs you are controlling | The script should start running. The Mixed Mode window will have a blinking Accept button. The Targets window will have a target of 1 for the DX of each LBCB. The test stats window should display step "1 0 4"
| Press Accept and then Start. While the ramp is running switch both Execute and Propose buttons to Auto. | The script will execute three steps. At the fourth step the script will send a close-session command to the OM after the OM finishes the ramp. At this point NTCP should disconnect and the Socket widget will disappear
| Type "testNtcpLink.pl --lbcb=? --option=5" in the command window where ? Is the number of LBCBs you are controlling | The script should start running. The Mixed Mode window will have a blinking Accept button. The Targets window will have a target of 1 for the DX of each LBCB. The test stats window should display step "1 0 4"
| Press Accept and then Start. While the ramp is running switch both Execute and Propose buttons to Auto. | The script will execute three steps. At the fourth step the script will send a close-session command to the OM after sending a get-control-point message. At this point NTCP should disconnect and the Socket widget will disappear
| Type "testNtcpLink.pl --lbcb=? --option=6" in the command window where ? Is the number of LBCBs you are controlling | The script should start running. The Mixed Mode window will have a blinking Accept button. The Targets window will have a target of 1 for the DX of each LBCB. The test stats window should display step "1 0 4"
| Press Accept and then Start. While the ramp is running switch both Execute and Propose buttons to Auto. | The script will execute three steps. At the fourth step the script will send a close-session command to the OM after the OM returns the control point data. At this point NTCP should disconnect and the Socket widget will disappear

## Oil Test with Free LBCB(s)
Test | Procedure | Result | Passed | Failed | Notes | 2012-03-21
Manual input complete step | Input a target of displacement Dx within the displacement limits and press "Set Target". Record the difference between the command displacement and the measured displacment as an offset. | Start button should blink
Input a ramp time of 10 seconds
| Press start. | The LBCB(s) should move for 10 seconds to the commanded displacement. The measured displacement should be the same as the commanded displacement when the offset is taken into account.
| Select 'Reset Targets To Zero' target pattern and click on 'Apply' | Target should consist of all 6 displacement DOFs set to zero. 'Start' button should blink.
| Press start. | All of the displacments should read zero
Repeat test for Dy
Repeat test for Dz
Repeat test for Rx
Repeat test for Ry
Repeat test for Rz
Repeat test for multiple DOFs and the other LBCB if available
Repeat test for the target patterns "Displacements Only", "Rotations Only", and "Displacements & Rotations"
Manual Input Global/Local toggles | Input a target of displacement Dx that is less than half way to the displacement limits and press "Set Target". | Start button should blink
Input a ramp time of 10 seconds
| Press start. | Mixed Mode State display should iterate through 'Ramp', 'Stab', and 'Hold'. Command at the end of the ramp should be the same as the target. The MM State should be 'Idle' at the end of the ramp. Record the resulting commands and displacements as 'target 1'
Open the Offsets window and set offsets for both command and measured displacements. Press OK to exit the window
| Click on the 'Global' button to switch it to 'Local' | Command and measured displacements should be extremely close to zero.
| Enter another target that is a small increment past zero and press "Set Target" | Start button should blink
| Press the start button | After the ramp is completed the command and displacement should display the small increment. The other displacements should still be zero. Record these commands and displacments as target 2.
| Switch from Local to Global | The displayed commands and displacements should equal target 2 + target 1.
| Enter a target that is a small increment past the current position and press "Set Target". | Start button should blink
| Press start | After the ramp is completed the command and displacement should display the change in displacement. Record these commands and displacements as target 3.
| Switch from Global to Local | The displayed commands and displacements should equal target 3 - target 1.
| Open the limits display and check the Local limit for a DOF. Record the global limit also. Close the limits display and set a target that is just beyond the local limit and press "Set Target". | Start button should blink
| Press Start | The ramp should run until the DOF reaches its limit.. At this point the MM state display should be 'Idle', The limit fault light should be red, and a message should be displayed that indicates which DOF has faulted.
| Switch from Local to Global | The command should be slightly beyond the global limit.
| Clear the fault by extending the limit. Switch from Global to Local | The fault light should turn green. The fault message should clear. The command for the faulted DOF should not change significantly.
| Select the reset to zero target pattern and press apply. | The start button should blink.
| Press Start | After the ramp runs, all commands and displacements should be zero.
| Swtich from Local to Global | The commands and displacements should equal target 1.
Repeat test for Dy
Repeat test for Dz
Repeat test for Rx
Repeat test for Ry
Repeat test for Rz
Repeat test for multiple DOFs and the other LBCB if available
Repeat test with LBCB 2 if testing with the two box setup
| Repeat test with the rigid connection button set to 'YES' if testing with the two box setup | In this case the offset widget will compute offsets for the combined LBCBs in addition to the separate boxes.

## Rubber Specimen Test
Test | Procedure | Result | Passed | Failed | Notes | 2012-04-04
Start | This test needs a rigid rubber specimen connected. The Startup procedure including oil turn on needs to done first. | The Mixed Mode widget should be ready to accept commands.
Force Control Test | Set the cartesian force limits to a window that is within the capacity of the rubber specimen. Move the LBCB a small amount in to translation and two rotation directions. The translation should be > 0.1 inches. The rotation should be > .001 radians. Make sure that none of the translations add tension to the specimen. | You should see a significant increase in force along the directions that are displaced. Record these force values as force target 1.
Enter ramp times of 2 seconds for all of the ramp entries in the Mixed Mode Window. Switch the LBCB Operation Manager Window to Mixed Mode. Set the Jacobian to 0.8 for all DOFs. Set the max increment to .01 for all translations and .0001 for all rotations. Set all Force Group Priorities to 5. Set the Force Error Range to 5 lbs for for the forces and 20 lbs*inches for the moments. Set the low stiffness limit to 20 and the high limit to 1E+6. Make sure that all of the force control parameters are set to "No".
| In the Mixed Mode Window set the forces moments which are in the same directions as the displacement DOFs set earlier to zero. Press Set Target then Start on the Mixed Mode Window. | In the LBCB Operation Manager Window you will see the Convergence display indicating which DOFs are converged as well as which DOF is being controlled. The stiffness matrix will have updated values for the column associated with the controlling DOF. The Target display will show for the controlling DOF a displacement that equals the current displacement plus the max increment. The last step will be slightly less than the max increment. The force ramp should end with all of the force DOFs converged. The resulting forces and moments should be 0 +/- the force error
| After the ramp is completed, set the force error range for one of the force DOFs to 3 lbs. Press Set Target and Start again. | The changed DOF will converge to the tighter force error range.
After the ramp is completed, set a limit for one of the target 1 forces to be just under the target 1 magnitude.
| Enter the forces for target 1 in the Mixed Mode Widget. Press Set Target and Start | The force ramp should stop on a limit fault for the limit that was changed
| After the ramp is stopped, set the Max increment for one of the force DOFs to .03 inches. Enter zero for all of the forces in the Mixed Mode Widget. Press Set Target and Start | Make sure that the force DOF with the changed Max Increment shows a displacement in the LBCB Operation Manager Target window that is the current displacement +/- .03.
| After the ramp is completed, reset the Max Increment to .01 inches. Set the Force Control Parameter "Use Force Error Range Ratios" to Yes. Enter the forces for target 1 and press Set Target and then Start | Take a look at which controlling DOF is chosen. In the previous ramps the moments were chosen first. In this case the force DOF with the 3 lb force error range should be chosen first.
| After the ramp is completed, set the Force Control Parameter "Use Force Error Range Ratios" to No. Set the Force Control Paraemter "Use Group Priorities" to Yes. Enter a different priority for each DOF in the Force Group Priorities display between 0 and 5. Enter zero for the forces and press Set Target and then Start | The controlling DOF that is chosen should follow the priorities with the lowest adjusted first.
| After the ramp is completed, set the Force Control Parameter "Multi Command Update" to Yes. Set Use Group Priorities to "NO". Set the Stiffness Test Ramp Count to 3. Enter the target 1 forces and press Set Target and then Start | First, all of the force DOFs that are not converged will be adjusted one by one to get new stiffness values. Then the OM will run 3 ramps with updated commands for all non-converged DOFs and no stiffness updates. After this the stiffness test is repeated
Force Control With Offsets Test | With the OM shut down and the oil off, use a tape measure to measure the difference between the top center of the rubber specimen and the center of the platten. Start the OM setup program and enter these measurements as the motion center for LBCB 1. Make sure that you account for the axis orientation of the LBCB | Since the rubber specimen is already connected, this relative tape measure is sufficient to determine the motion center. However for the outside tests, the motion center needs to be measured before connection so that the krypton alignment is referenced to this control point. The cartesian position of the LBCB needs to be include in the calculation of the motion center because of this.
| Use manual force commands to minimize all of the forces to as close to zero as possible. Take a snapshot after the minimization is completed and label it Run 1 | The forces will be close to zero. However there will be some residual displacement readings due to deformations in the rubber.
Open the Offsets widget and set the displacement command offsets and displacement offsets to the current position using the "Execute" feature.
| Toggle the Global button to Local and back several times and not how all of the displacements change in both the LBCB Operations Manager window and the Mixed Mode window. | The displacements in the Mixed Mode widget and the LBCB Operations Manager should read zero when the "Global" button is set to "Local". This is true for both cartesian and actuator space displacement commands and readings.
In the Offsets Widget, create a new force offset at the current position. Call this offset "Test1"
| Set the force tag to "Test1" and toggle the Global button to Local and back several times and not how all of the numbers change in both the LBCB Operations Manager window and the Mixed Mode window. | Both displacements and forces in the Mixed Mode widget and the LBCB Operations Manager should read zero when the "Global" button is set to "Local". This is true for both cartesian and actuator space displacement commands and readings.
| Switch back to Local. In the Mixed Mode window select Dx,Dz,Rz, Fz, and My targets. Toggle the Local button to Global. | All of the targets should be set to the offset values with the signs reversed.
| Switch back to Local. Clear the targets. Apply axial force to the specimen by setting Fz to between 10 and 20 pounds | After the force command is finished, Fz should read whatever was targeted.
Open the Offset widget and set another Force offset with the tag "Test2".
| Switch the Force tag on the Mixed Mode widget to "Test 2" | Fz should be zero when in Local
Shut down the oil and restart the OM
| After autobalancing, open the Offsets window and load the Offsets from Run 1. | In the Offsets window when the Force offsets are selected, The force tag selector should contain <None>, Test, and Test2
Switch the Force tag on the Mixed Mode widget to "Test 1". Switch to Local
| In the Mixed Mode window, Select the "Reset to Zero" pattern and execute this target. | After the command is finished all forces and displacements should be zero.
| Switch the force tag in the Mixed Mode window to "Test 2" | The Fz command should read -10 to -20 pounds depending on what axial force you applied earlier
Shut down the oil and start up OM Setup. Set the motion center to all zeros. Restart the OM, autobalance and turn on the oil.
Load the offsets from Run 1
Switch the Force tag on the Mixed Mode widget to "Test 1". Switch to Local
| In the Mixed Mode window, Select the "Reset to Zero" pattern and execute this target. | All of the commands and displacments in actuator space should be zero. The displacements and forces in cartesian space will not necessarily be zero because of the change in geometry due to the new motion center.
Take another snapshot and call it Run 2.
| Compare the actuator displacments and forces between run 1 and run 2 | All of the values will be the same. The cartesian values will be different due to the change in motion center. For extra credit you could do a force equilibrium analysis with the actuator forces to determine if the change in cartesian values are correct

## Fake Startup
Test | Procedure | Result | Passed | Failed | 2012-03-21
Fake Startup | Double click on the file "LBCB Operation Manager Project.lvproj" | LabVIEW should start up and display the project tree
| Double click on "LBCB Operations Manager.vi" and "Mixed Mode Widget.vi" | Both vi's should open with a clean arrow
| Press the blue Real button | The button should turn pink and say Fake
| Press the start arrow on the LBCB Operation Manager and then the Mixed Mode Widget | Both widgets should be running. The Mixed Mode target and input mode window should be grayed out.
| Press Start Autobalancing on the LBCB Operation Manager | The auto balance initialization and Command Displacement Difference window should appear
| Select autobalance on all auto balance windows that appear. | A warning popup should show up. After ok'ing the popup, the system should start autobalancing. After it is balanced there will most likely be numerous limit faults because the fake initial position is deliberately skewed. The limits will need to be adjusted to correct these faults before any further testing is done.
| Press control on one or both LBCBs | The Mixed Mode Widget should now allow inputs
Real Startup | Open up the OM Latest Build folder and double click on the "LBCB Operation Manager" executable | Two widgets should be running. For the Mixed Mode widget, target and input mode window should be grayed out.
| Press Start Autobalancing on the LBCB Operation Manager | The auto balance initialization and Command Displacement Difference window should appear
| Select autobalance on all auto balance windows that appear. | A warning popup should show up. After ok'ing the popup, the system should start autobalancing. After it is balanced there will most likely be numerous limit faults because the fake initial position is deliberately skewed. The limits will need to be adjusted to correct these faults before any further testing is done.
| Press control on one or both LBCBs | The Mixed Mode Widget should now allow inputs
| Check the difference between command and displacement for all of the actuators and DOFs. | The differences should be insignificant. If they are not the LBCB will jump to its initial position once oil pressure is applied
| For the small LBCBs turn low pressure on. | Verify that the forces have not changed significantly
| For the large LBCBs turn low pressure on and unlock the actuators one at a time starting with the Z actuators. | Verify that the forces have not changed significantly
| Switch to high pressure | The forces will jump somewhat but then settle to an equilibrium. This equilibrium is zero unless the LBCB is connected to a specimen. If a specimen is connected then the equilbrium will be close to the last position the LBCB was moved to before it was shut down.