## LBCB Serial Numbers
LBCB BOX -Load Cell Serial Numbers
Slope
141728A | X1 | 32.952
141801A | X2 | 33.1353
141730A | Y1 | 34.0686
141806A | Z1 | 34.4601
141803A | Z2 | 33.4389
141802A | Z3 | 32.7226
141801A | X2
141802A | Z3
141803A | Z2
141806A | Z1
141728A | X1
141730A | Y1

## X1 LOAD CELL
FACTORY CALIBRATION
Load Cell Capacity | KIP | Lbs | N
| 270 | 270000 | 1201019.832
Bridge Resistance | Ω
350.8
Shunt Resistance | KΩ | Ω
| 30 | 30000
Excitation | V
10
Rated Output | Positive mV/V | Negative mV/V
| 4.1519 | -4.1372
Zero Balance | %RO
-0.051
Shunt Output | Positive mV/V | Negative mV/V
| 2.9022 | -2.9014
Shunt Load Positive | KIP | Lbs | N
| 188.88 | 188880 | 840180.0958
Shunt Load Negative | KIP | Lbs | N
| -189.27 | -189270 | -841914.9022
Factory Slope | KIP/mV/V | Lbs/mV/V | N/mV/V
| 65.1458 | 65145.7939 | 289782.9275
Factory Offset | KIP | Lbs | N
| -0.1377 | -137.7 | -612.5201
INSTALLED CALIBRATION
Shunt Resistance | KΩ | Ω
| 60 | 60000
Excitation | V
Shunt Output | Positive mV/V | Negative mV/V
| 1.4595 | -1.4591
Shunt Load Positive | KIP | Lbs | N
| 94.9449 | 94944.8521 | 422335.7419
Shunt Load Negative | KIP | Lbs | N
| -95.194 | -95194.0423 | -423444.1952
MEASURED VOLTAGES
| V | KIP | Lbs | N
Shunt Positive | 3.0487 | 94.9449 | 94944.8521 | 422335.7419
Zero Load | 0.16232 | 0 | 0 | 0
Shunt Negative | -2.7215 | -95.194 | -95194.0423 | -423444.1952
[5 empty] | Slope | 32.952
[5 empty] | R2 | 1
[5 empty] | Voltage @ Max Kip | 10.0146
[5 empty] | Max Kip Compression | 334.8685
[5 empty] | Max Kip Tension | -324.1711
[5 empty] | Previous Slope | 16.9175

## X2 LOAD CELL
FACTORY CALIBRATION
Load Cell Capacity | KIP | Lbs | N
| 270 | 270000 | 1201019.832
Bridge Resistance | Ω
351.1
Shunt Resistance | KΩ | Ω
| 30 | 30000
Excitation | V
10
Rated Output | Positive mV/V | Negative mV/V
| 4.1405 | -4.13
Zero Balance | %RO
0.036
Shunt Output | Positive mV/V | Negative mV/V
| 2.9143 | -2.909
Shunt Load Positive | KIP | Lbs | N
| 190.16 | 190160 | 845873.8195
Shunt Load Negative | KIP | Lbs | N
| -190.15 | -190150 | -845829.3372
Factory Slope | KIP/mV/V | Lbs/mV/V | N/mV/V
| 65.2923 | 65292.304 | 290434.6368
Factory Offset | KIP | Lbs | N
| 0.0972 | 97.2 | 432.3671
INSTALLED CALIBRATION
Shunt Resistance | KΩ | Ω
| 60 | 60000
Excitation | V
Shunt Output | Positive mV/V | Negative mV/V
| 1.4656 | -1.463
Shunt Load Positive | KIP | Lbs | N
| 95.7914 | 95791.3734 | 426101.2564
Shunt Load Negative | KIP | Lbs | N
| -95.4229 | -95422.9422 | -424462.3928
MEASURED VOLTAGES
| V | KIP | Lbs | N
Shunt Positive | 2.9597 | 95.7914 | 95791.3734 | 426101.2564
Zero Load | 0.07088 | 0 | 0 | 0
Shunt Negative | -2.811 | -95.4229 | -95422.9422 | -424462.3928
[5 empty] | Slope | 33.1353
[5 empty] | R2 | 1
[5 empty] | Voltage @ Max Kip | 9.9592
[5 empty] | Max Kip Compression | 333.7013
[5 empty] | Max Kip Tension | -329.0041
[5 empty] | Previous Slope | 16.4798

## Y1 LOAD CELL
FACTORY CALIBRATION
Load Cell Capacity | KIP | Lbs | N
| 270 | 270000 | 1201019.832
Bridge Resistance | Ω
350.1
Shunt Resistance | KΩ | Ω
| 30 | 30000
Excitation | V
10
Rated Output | Positive mV/V | Negative mV/V
| 4.1119 | -4.1048
Zero Balance | %RO
0.141
Shunt Output | Positive mV/V | Negative mV/V
| 2.9025 | -2.9007
Shunt Load Positive | KIP | Lbs | N
| 190.59 | 190590 | 847786.5547
Shunt Load Negative | KIP | Lbs | N
| -190.79 | -190790 | -848676.1991
Factory Slope | KIP/mV/V | Lbs/mV/V | N/mV/V
| 65.7198 | 65719.8145 | 292336.2985
Factory Offset | KIP | Lbs | N
| 0.3807 | 380.7 | 1693.438
INSTALLED CALIBRATION
Shunt Resistance | KΩ | Ω
| 60 | 60000
Excitation | V
Shunt Output | Positive mV/V | Negative mV/V
| 1.4597 | -1.4588
Shunt Load Positive | KIP | Lbs | N
| 96.3099 | 96309.8706 | 428407.6469
Shunt Load Negative | KIP | Lbs | N
| -95.489 | -95488.9797 | -424756.142
MEASURED VOLTAGES
| V | KIP | Lbs | N
Shunt Positive | 2.9923 | 96.3099 | 96309.8706 | 428407.6469
Zero Load | 0.17556 | 0 | 0 | 0
Shunt Negative | -2.6374 | -95.489 | -95488.9797 | -424756.142
[5 empty] | Slope | 34.0686
[5 empty] | R2 | 1
[5 empty] | Voltage @ Max Kip | 9.6863
[5 empty] | Max Kip Compression | 346.667
[5 empty] | Max Kip Tension | -334.7047
[5 empty] | Previous Slope | 16.6121

## Z1 LOAD CELL
FACTORY CALIBRATION
Load Cell Capacity | KIP | Lbs | N
| 270 | 270000 | 1201019.832
Bridge Resistance | Ω
350.7
Shunt Resistance | KΩ | Ω
| 30 | 30000
Excitation | V
10
Rated Output | Positive mV/V | Negative mV/V
| 4.0941 | -4.0771
Zero Balance | %RO
0.125
Shunt Output | Positive mV/V | Negative mV/V
| 2.9023 | -2.903
Shunt Load Positive | KIP | Lbs | N
| 191.57 | 191570 | 852145.8119
Shunt Load Negative | KIP | Lbs | N
| -192.18 | -192180 | -854859.2271
Factory Slope | KIP/mV/V | Lbs/mV/V | N/mV/V
| 66.0858 | 66085.7646 | 293964.1257
Factory Offset | KIP | Lbs | N
| 0.3375 | 337.5 | 1501.2748
INSTALLED CALIBRATION
Shunt Resistance | KΩ | Ω
| 60 | 60000
Excitation | V
Shunt Output | Positive mV/V | Negative mV/V
| 1.4596 | -1.4599
Shunt Load Positive | KIP | Lbs | N
| 96.7951 | 96795.1376 | 430566.222
Shunt Load Negative | KIP | Lbs | N
| -96.1434 | -96143.4021 | -427667.1578
MEASURED VOLTAGES
| V | KIP | Lbs | N
Shunt Positive | 3.2391 | 96.7951 | 96795.1376 | 430566.222
Zero Load | 0.44022 | 0 | 0 | 0
Shunt Negative | -2.3598 | -96.1434 | -96143.4021 | -427667.1578
[5 empty] | Slope | 34.4601
[5 empty] | R2 | 1
[5 empty] | Voltage @ Max Kip | 9.5763
[5 empty] | Max Kip Compression | 359.7707
[5 empty] | Max Kip Tension | -329.4308
[5 empty] | Previous Slope | 16.8241

## Z2 LOAD CELL
FACTORY CALIBRATION
Load Cell Capacity | KIP | Lbs | N
| 270 | 270000 | 1201019.832
Bridge Resistance | Ω
350.2
Shunt Resistance | KΩ | Ω
| 30 | 30000
Excitation | V
10
Rated Output | Positive mV/V | Negative mV/V
| 4.1221 | -4.1118
Zero Balance | %RO
-0.07
Shunt Output | Positive mV/V | Negative mV/V
| 2.9046 | -2.9009
Shunt Load Positive | KIP | Lbs | N
| 190.42 | 190420 | 847030.3571
Shunt Load Negative | KIP | Lbs | N
| -190.44 | -190440 | -847119.3215
Factory Slope | KIP/mV/V | Lbs/mV/V | N/mV/V
| 65.5825 | 65582.5308 | 291725.6299
Factory Offset | KIP | Lbs | N
| -0.189 | -189 | -840.7139
INSTALLED CALIBRATION
Shunt Resistance | KΩ | Ω
| 60 | 60000
Excitation | V
Shunt Output | Positive mV/V | Negative mV/V
| 1.4607 | -1.4589
Shunt Load Positive | KIP | Lbs | N
| 95.6092 | 95609.1998 | 425290.9079
Shunt Load Negative | KIP | Lbs | N
| -95.8652 | -95865.1681 | -426429.5115
MEASURED VOLTAGES
| V | KIP | Lbs | N
Shunt Positive | 3.1474 | 95.6092 | 95609.1998 | 425290.9079
Zero Load | 0.28216 | 0 | 0 | 0
Shunt Negative | -2.5787 | -95.8652 | -95865.1681 | -426429.5115
[5 empty] | Slope | 33.4389
[5 empty] | R2 | 1
[5 empty] | Voltage @ Max Kip | 9.8688
[5 empty] | Max Kip Compression | 343.8238
[5 empty] | Max Kip Tension | -324.9534
[5 empty] | Previous Slope | 16.6052

## Z3 LOAD CELL
FACTORY CALIBRATION
Load Cell Capacity | KIP | Lbs | N
| 270 | 270000 | 1201019.832
Bridge Resistance | Ω
350.8
Shunt Resistance | KΩ | Ω
| 30 | 30000
Excitation | V
10
Rated Output | Positive mV/V | Negative mV/V
| 4.1517 | -4.1434
Zero Balance | %RO
0.166
Shunt Output | Positive mV/V | Negative mV/V
| 2.9028 | -2.9026
Shunt Load Positive | KIP | Lbs | N
| 188.9 | 188900 | 840269.0602
Shunt Load Negative | KIP | Lbs | N
| -189.1 | -189100 | -841158.7046
Factory Slope | KIP/mV/V | Lbs/mV/V | N/mV/V
| 65.0987 | 65098.6727 | 289573.3221
Factory Offset | KIP | Lbs | N
| 0.4482 | 448.2 | 1993.6929
INSTALLED CALIBRATION
Shunt Resistance | KΩ | Ω
| 60 | 60000
Excitation | V
Shunt Output | Positive mV/V | Negative mV/V
| 1.4598 | -1.4597
Shunt Load Positive | KIP | Lbs | N
| 95.4816 | 95481.6202 | 424723.4056
Shunt Load Negative | KIP | Lbs | N
| -94.5787 | -94578.6725 | -420706.8941
MEASURED VOLTAGES
| V | KIP | Lbs | N
Shunt Positive | 3.1657 | 95.4816 | 95481.6202 | 424723.4056
Zero Load | 0.26141 | 0 | 0 | 0
Shunt Negative | -2.6425 | -94.5787 | -94578.6725 | -420706.8941
[5 empty] | Slope | 32.7226
[5 empty] | R2 | 0.99999
[5 empty] | Voltage @ Max Kip | 10.0848
[5 empty] | Max Kip Compression | 335.7804
[5 empty] | Max Kip Tension | -318.6725
[5 empty] | Previous Slope | 16.5706