## Displacement Cals
X1 |  |  | X2
Cmd Volts | LVDT Reading Volts | Caliper Reading inches | Cmd Volts | LVDT Reading Volts | Caliper Reading inches
0 | -0.088761 | 3.6275 | 0 | 0.033558 | 3.6205
1 | -1.1056 | 3.353 | 1 | -0.97032 | 3.349
2 | -2.1156 | 3.083 | 2 | -1.9736 | 3.079
3 | -3.1256 | 2.8105 | 3 | -2.9762 | 2.8085
2 | -2.1103 | 3.082 | 2 | -1.965 | 3.0775
1 | -1.0989 | 3.355 | 1 | -0.96166 | 3.354
0 | -0.087902 | 3.6275 | 0 | 0.042007 | 3.6225
-1 | 0.92307 | 3.8945 | -1 | 1.0458 | 3.896
-2 | 1.9343 | 4.1755 | -2 | 2.0413 | 4.1715
-3 | 2.9453 | 4.449 | -3 | 3.0537 | 4.4415
-2 | 1.9286 | 4.174 | -2 | 2.0414 | 4.1655
-1 | 0.91712 | 3.9 | -1 | 1.0373 | 3.895
0 | -0.094363 | 3.627 | 0 | 0.032881 | 3.624
New Slope: | -0.27292 | 0.26983 | New Slope: | -0.27232 | 0.27126
| in/cmd | in/LVDT |  | in/cmd | in/LVDT
R-squared value | 0.99998 | 0.99998 | R-squared value | 0.99998 | 0.99998
Previous Slopes: | -1.9572 | 1.9242 | Previous Slopes: | -1.9705 | 1.9512
Delta | 1.6843 | -1.6544 | Delta | 1.6982 | -1.68
% Difference | 151% | 150.8% | % Difference | 151.4% | 151.2%
Y1 |  |  | Z1
Cmd Volts | LVDT Reading Volts | Caliper Reading inches | Cmd Volts | LVDT Reading Volts | Caliper Reading inches
0 | -0.064016 | 2.5835 | 0 | -0.080935 | 2.6235
1 | -1.0555 | 2.4595 | 1 | -1.0805 | 2.4835
2 | -2.0501 | 2.308 | 2 | -2.0811 | 2.348
3 | -3.0447 | 2.1855 | 3 | -3.0812 | 2.2145
2 | -2.0501 | 2.312 | 2 | -2.0828 | 2.352
1 | -1.0555 | 2.4495 | 1 | -1.082 | 2.4845
0 | -0.060347 | 2.58 | 0 | -0.08094 | 2.6155
-1 | 0.93529 | 2.7185 | -1 | 0.92014 | 2.753
-2 | 1.9305 | 2.85 | -2 | 1.9218 | 2.8875
-3 | 2.9265 | 2.9865 | -3 | 2.9231 | 3.021
-2 | 1.9304 | 2.8485 | -2 | 1.9231 | 2.887
-1 | 0.93513 | 2.7135 | -1 | 0.92149 | 2.754
0 | -0.059756 | 2.585 | 0 | -0.079628 | 2.616
New Slope: | -0.13376 | 0.13441 | New Slope: | -0.13441 | 0.13429
| in/cmd | in/LVDT |  | in/cmd | in/LVDT
R-squared value | 0.99967 | 0.99967 | R-squared value | 0.99993 | 0.99992
Previous Slopes: | -0.98295 | 0.97597 | Previous Slopes: | -0.9964 | 0.99664
Delta | 0.84919 | -0.84156 | Delta | 0.86199 | -0.86235
% Difference | 152.1% | 151.6% | % Difference | 152.5% | 152.5%
Z2 |  |  | Z3
Cmd Volts | LVDT Reading Volts | Caliper Reading inches | Cmd Volts | LVDT Reading Volts | Caliper Reading inches
0 | 0.19787 | 2.605 | 0 | -0.20094 | 2.621
1 | -0.80719 | 2.4735 | 1 | -1.2005 | 2.4895
2 | -1.812 | 2.338 | 2 | -2.2012 | 2.359
3 | -2.8167 | 2.204 | 3 | -3.2018 | 2.2285
2 | -1.8109 | 2.372 | 2 | -2.1994 | 2.3645
1 | -0.80574 | 2.484 | 1 | -1.1995 | 2.5005
0 | 0.19971 | 2.607 | 0 | -0.19903 | 2.621
-1 | 1.2053 | 2.7405 | -1 | 0.80201 | 2.7515
-2 | 2.2115 | 2.873 | -2 | 1.8046 | 2.882
-3 | 3.2175 | 3.007 | -3 | 2.8051 | 3.0145
-2 | 2.211 | 2.874 | -2 | 1.8029 | 2.8835
-1 | 1.2048 | 2.7395 | -1 | 0.80078 | 2.7535
0 | 0.19901 | 2.606 | 0 | -0.19936 | 2.6225
New Slope: | -0.13172 | 0.13098 | New Slope: | -0.13045 | 0.13031
| in/cmd | in/LVDT |  | in/cmd | in/LVDT
R-squared value | 0.99856 | 0.99857 | R-squared value | 0.99983 | 0.99983
Previous Slopes: | -0.98256 | 0.97106 | Previous Slopes: | -0.9738 | 0.96773
Delta | 0.85083 | -0.84008 | Delta | 0.84335 | -0.83742
% Difference | 152.7% | 152.5% | % Difference | 152.7% | 152.5%

## SW Cmd Cals
| Full Extension |  | Full Retraction |  | Full Retraction After Y1 Calibration |  |  | Mid-Point
| %FS | LVDT Volts | %FS | LVDT Volts | %FS | LVDT Volts | Inches | %FS | LVDT Volts | Inches
X1 | -70.11 | 7.363 | 72.09 | -7.558
X2 | -72.01 | 7.462 | 71.08 | -7.386
Y1 | -71.03 | 7.316 | 72.72 | -7.453 | 72.72 | -7.453
Z1 |  |  | 73.22 | -7.526 | 73.22 | -7.526
Z2 |  |  | 70.64 | -7.296 | 70.64 | -7.296
Z3 |  |  | 76.49 | -7.891 | 76.49 | -7.891
The concept is to set up a SW command calibration that aligns with the mechanical limits of the LBCB so that the zero volt mid-stroke is the came as mechanical mid-stroke.
1. Set the SW command calibration to the default.
2. Locate the mechanical limits of the actuator in %FS.
3. Set a new command calibration with the two limit points.
For this particular calibration, we have a specimen in the way that prevents us from fully extending the Z actuators. For this reason we measure the full retracted and mid-point positions of the Y actuator and then use those measurements to establish the mechanical mid-point of the Z actuators.